## Slide 1: Unexpected Uses of 3D Printing During COVID-19: The Creation and Clinical Validation of a 3D Printed Nasopharyngeal Swab for Critical Need Testing
Todd Goldstein PhD, Jonathan Ford PhD, Sean Trahan MS, Todd Hazelton MD, and Summer Decker PhD
Northwell Health Systems and USF Health Morsani College of Medicine, Tampa General Hospital

## Slide 2: COVID-19 Testing Kit & Crisis
Kit consists of three key parts
  Test Tube
  Viral Transport Media
  Nasopharyngeal Swab
  Synthetic flocked swabs standard as organic (cotton) swabs interfere with viral testing
  Global supply chain has been disrupted directly impacting the ability to test in the US
  (Main producers are Italy and China)

## Slide 3: 3D Printed Swabs
USF Radiology Division of 3D Clinical Applications
  Point-of-Care 3D Printing lab at Tampa General Hospital
Northwell Health’s Division of 3D Design and Innovation- New York City

## Slide 4: Design
Design was a collaborative process
We designed several tip and stem iterations
Received feedback from clinicians
Tested on volunteers for ergonomic comfort
Finalized design before bench lab testing

## Slide 5: Printer and Material Selection
3D Printer Type and Material
  Formlabs Form 2 and Form 3B
  Surgical Guide Resin - Already cleared by FDA for clinical and dental use
  Sterilizable via autoclave
  Existing widespread use in Clinical and Dental 3D Printing so already in many hospitals and clinics

## Slide 6: Volunteer Testing
Tested comfort and ergonomics on volunteers
Notes: Jonathan swabbing his head

## Slide 7: Bench Lab Testing: Can it detect viral loads?
1 ml of synthetic sputum (from Sophie Darch, Mol Med) was spiked or not with 106 PFU of RSV.
Swabs were dipped into the sputum, rotated for 5 sec, then the tips were cut off into BD UTM (1 ml).
Tips were incubated in UTM for 15 min, 22oC
140 ml UTM from each tube was subjected to RNA isolation with Qiagen viral RNA mini kit, elution in 50 ml.
10 ul of RNA was used to make cDNA (iScript, BioRad) in a 20 ml final volume.
qRT-PCR was performed with 2 sets of RSV primers for the N gene and the SH-G intergenic region. 2 ml of cDNA per well, triplicate wells. 20 ml qRT-PCR using SensiFAST SYBRgreen kit (Bioline).
Shown are cycle thresholds (C(t)), where RSV RNA was detected. Lower is better.
Conclusion: 3DP Swab statistically just as good or better than current synthetic swab
Chart: RSV N
| RSV | No RSV
Decker | 5.8868 | 23.1906
Spun polyester | 9.0649 | 23.3479
C(t)
Chart: RSV SH-G
| RSV | No RSV
Decker | 7.2336 | 26.2379
Spun polyester | 8.0366 | 26.1453
C(t)
Notes: Dr. Michael Teng

## Slide 8: Will it interfere with viral testing assay?
2 ml of BD UTM was spiked or not with 106 PFU of RSV and aliquoted to 2 1.8 ml cryotubes.
One sterile swab tip was cut off into one tube.
Cryotubes were incubated at 4oC, 24, 48 and 72 hrs.
140 ml UTM from each tube was subjected to RNA isolation with Qiagen viral RNA mini kit, elution in 50 ml.
10 ul of RNA was used to make cDNA (iScript, BioRad) in a 20 ml final volume.
qRT-PCR was performed with 2 sets of RSV primers for the N gene and the SH-G intergenic region. 2 ml of cDNA per well, triplicate wells. 20 ml qRT-PCR using SensiFAST SYBRgreen kit (Bioline).
Shown are cycle thresholds (C(t)), where RSV RNA was detected.
Conclusion: 3DP Swabs do not leech chemicals that interfere with viral assay
Chart: Genome
| Day 1 | Day 2 | Day 3
No swab | 27.3927 | 27.561 | 26.4834
EtOH | 27.0977 | 27.1902 | 26.9603
Autoclave | 27.1713 | 27.3789 | 26.8665
Chart: N
| Day 1 | Day 2 | Day 3
No swab | 24.7881 | 25.4079 | 24.5937
EtOH | 24.7998 | 25.3102 | 25.4023
Autoclave | 24.7986 | 25.4227 | 24.9579

## Slide 9
USF Health/Tampa General Hospital
Northwell Health
Thomas Jefferson University
Direct head-to-head comparison on same patient in alternative nostrils
  Synthetic swab in one nostril
  3DP swab in alternate nostril
Results of 291 patients
  156 Females, 135 Males
Significance
  3 different assay machines
Clinical Trial Testing: Comparison of C(t) for Viral Targets
Conclusion: Cycle threshold was statistically the same so the 3DP swab as good as the synthetic standard, in some cases better.
Hospital review boards accepted 3DP swab as alternative swab.
 |  | 3DP | 3DP | 3DP | Total
 |  | + | - | inc | 
FLNP | + | 74 | 2 | 3 | 79
FLNP | - | 4 | 203 | 1 | 208
FLNP | inc | 0 | 2 | 2 | 4
Total |  | 78 | 207 | 6 | 291

## Slide 10: But does it hurt?
3DP swabs performed similarly to standard of care swabs with respect to patient comfort

## Slide 11: Completing the Kit
Swabs are delivered to USF Health Dept. of Infectious Disease labs for sterilization
Each swab is individually wrapped before autoclaving
Bundled with the other testing kit components
Distributed to affiliate hospitals

## Slide 12: Global Impact
US VA system- 25% of the 170 hospitals are using the USF Health/Northwell Health swab
  Majority are printing their own; some are being supplied by affiliates like USF MCOM
  Doubled the 3D printing in VA facilities across the US
Numerous states using the USF Health/Northwell Health swab including New York, Massachusetts, Kansas, Oklahoma, and Ohio
  Ohio has over 1 million alone
Over 60 hospitals and hospital systems including the Mayo Clinic, Brigham and Women’s, LSU Medical Center, etc.
Used in 30+ countries: Canada, Colombia, Brazil, Japan, Switzerland, Ethiopia, Malawi, Uruguay, Georgia, Portugal, Malaysia, Argentina, United Kingdom, etc.
In total, there are tens of millions USF Health/Northwell Health designed swabs in the world which have directly impacted the national response to COVID.

## Slide 13: 3D Printed NP Swab Team
USF Health
  Dr. Summer Decker*
  Dr. Jonathan Ford*
  Dr. Todd Hazelton
  Dr. Michael Teng
  Dr. Kami Kim
  Dr. Jason Wilson
  Dr. John Sinnott
  Dr. Suzane Silbert
  Dr. Charles Lockwood
Northwell Health
  Dr. Todd Goldstein*
  Dr. Greg Berry
  Dr. Ona Bloom
  Dr. James Crawford
  Dr. Dwayne Breining
  Mr. James Castagnaro
Thomas Jefferson University
  Dr. Robert Pugliese
  Dr. Matthew Pettengill
  Dr. Kristy Shine
  Dr. Zi-Xuan Wang
  Dr. Morgan Hutchinson
  Dr. Barbara Goldsmith
  Dr. Donna Gentile O’Donnell

## Slide 14: Thank you
For more information on swabs please contact:
3dclinicalapplications@usf.edu